Daily Response Check Sheet

Detector Model: _____________________        		Detector Serial #:

Manufacturer’s Last Calibration Date:

DO NOT USE ANY TESTING EQUIPMENT WITHOUT VERIFYING CALIBRATION.

1.	ZERO THE INSTRUMENT IN FRESH AIR

2.	CHECK BATTERY.

3.	CHECK INTERNAL PUMP (If Fitted).

4.	EMPTY AND CLEAN LINE TRAP (If Fitted).

5.	VERIFY CALIBRATION O2 METER AT 20.8% or 20.9% (Adjust if necessary)

6.	VERIFY COMBUSTIBLE METER ZERO (Adjust if necessary).

7.	ZERO H2S METER (If Fitted).

8.	RESPONSE TEST COMBUSTIBLE METER WITH TEST GAS. (Ensure reading within target range)

9.	CHECK ALARM(S) (If Fitted).

10.	CHECK LIGHT (If Fitted).

11.	RESPONSE CHECK H2S METER WITH TEST GAS (If Fitted and H2S Test is required)

12.	RECORD SIGNATURE (INITIALS) IN RESPONSE CHECK TABLE BELOW

NOTE:  Testing equipment that fails any check must be immediately taken out of service.  Affix a “Danger – Do Not Operate” tag, and contact the Manufacturer’s Representative to arrange equipment service.

| Year: 20...... | Jan | Feb | Mar | Apr | May | Jun | Jul | Aug | Sep | Oct | Nov | Dec |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 1 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 2 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 3 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 4 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 5 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 6 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 7 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 8 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 9 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 10 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 11 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 12 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 13 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 14 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 15 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 16 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 17 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 18 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 19 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 20 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 21 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 22 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 23 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 24 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 25 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 26 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 27 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 28 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 29 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 30 |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 31 |  |  |  |  |  |  |  |  |  |  |  |  |

# Text boxes

DAY